## Summary
Summary
Drug | Cocaine | [6 empty] | Shorthand |  | Coc
ISTD | Cocaine-d8 | [8 empty] | Coc-d8
Calibration range: | 2.5 | 5 | 25 | 50 | 75 | 100 | 150 | 200 | µg/L
QCs: | 10 | 50 | 140 | µg/L
Specified Limit: | 10 | µg/L
[6 empty] | QC 1 |  | QC 2 |  | QC 3
Concentration (µg/L) | [5 empty] | 10 |  | 50 |  | 140
Bias (%) | [5 empty] | -12.0397 |  | -11.3692 |  | -5.7907
Accuracy (%) | [5 empty] | 87.9603 |  | 88.6308 |  | 94.2093
% CV | [5 empty] | 7.4966 |  | 5.0294 |  | 4.9331
Method Uncertainty Contributions
Precision | [5 empty] | 0.031976 |  | 0.015222 |  | 0.011979
Standard solution prep & Spiking | [5 empty] | 0.040347 |  | 0.040347 |  | 0.040347
Sample matrix volume | [5 empty] | 0.0052 |  | 0.0052 |  | 0.0052
Calibration curve | [5 empty] | 0.054214 |  | 0.05439 |  | 0.054422
Analytical measurement | [5 empty] | 0.06758 |  | 0.067721 |  | 0.067747
Combined uncertainty | [5 empty] | 0.74944 |  | 3.4803 |  | 9.6592
Expanded uncertainty (99.7 % confidence interval) | [5 empty] | 2.2483 |  | 10.4408 |  | 28.9775
Expanded uncertainty (%) | [5 empty] | 22.4831 |  | 20.8815 |  | 20.6982
Forensic Science Regulator set UM at Specified Limit | [5 empty] | 35%

## Note
Note
Quantifying the Sources of Uncertainity
Uncertainty of Precision
The precision uncertainty tab calculates different values for each nominal value used.
Use the precsion value which as a nominal concentration closest to the mean of the analytical results obtained for the case sample.
Uncertainty of Standard Solution Preparation and Calibration Curve Spiking
To determine the uncertainty of the standard solution(s) and calibration curve spiking, the following parameters for each standard solution and calibrator need to be entered:
1. The purity and accuracy of the reference standard (CRM).
2. The volume, accuracy and number of times each pipette and/or volumetric flask is used.
Uncertainty of Calibration and Analytical Measurement
Identify the calibration curve with the largest standard error (SE). The calibration data of that curve ('Calibration & Measurement' cells C30-D39) is used to calculate the regression parameters.
Computing the uncertainty of calibration requires specifiying the mean analytical result of a case sample, and the number of replicates used to determine that concentration.
The uncertainty of calibration (and analytical measurement) can then be computed using the specified formulae. The computations have been demonstrated for duplicate results at the 3 QC levels.
Recommendations for reducing measurement uncertainty
1. Narrow the calibration concentration range around the specified limit concentration.
2. Add more data.

## Raw Data
Accuracy Data | - Cocaine
QC 1 | Nominal Conc (µg/L) | Run 1 |  | Run 2 |  | Run 3 |  | Run 4 |  | Run 5 |  | Run 6 |  | Run 7 |  | Run 8 |  | Run 9 |  | Run 10 |  | Run 11
|  | 2019-07-26 |  | 2019-07-29 |  | 2019-07-30 |  | 2019-07-31 |  | 2019-08-01 |  | 2019-08-02 |  | 2019-08-05 |  | 2019-08-06 |  | 2019-08-07 |  | 2019-08-08 |  | 2019-08-09
Rep 1 | 10 | 12 | 10.0877 | 9.2787 | 9.2679 | 9.0924 | 9.1532 | 10.1085 | 10.0821 | 8.8837 | 8.9801 | 7.3419 | 7.4109 | 8.0019 | 8.0344 | 8.7642 | 8.6388 | 8.3954 | 8.3145 | 8.5137 | 8.5404 | 9.2988 | 9.3299
|  | 8.1753 |  | 9.2571 |  | 9.214 |  | 10.0557 |  | 9.0765 |  | 7.4799 |  | 8.0669 |  | 8.5135 |  | 8.2335 |  | 8.5671 |  | 9.361
Rep 2 |  | 8.6618 | 8.5351 | 9.2985 | 9.3357 | 9.3141 | 9.2757 | 8.866 | 8.9436 | 8.4763 | 8.5131 | 7.849 | 8.014 | 7.9323 | 8.014 | 8.4685 | 8.4105 | 8.6039 | 8.602 | 8.547 | 8.7667 | 9.1131 | 9.2625
|  | 8.4084 |  | 9.3729 |  | 9.2373 |  | 9.0212 |  | 8.5499 |  | 8.1789 |  | 8.0958 |  | 8.3526 |  | 8.6 |  | 8.9864 |  | 9.4118
Mean |  |  | 9.3114 |  | 9.3018 |  | 9.2144 |  | 9.5129 |  | 8.7466 |  | 7.7124 |  | 8.0242 |  | 8.5247 |  | 8.4582 |  | 8.6535 |  | 9.2962
Std Dev |  |  | 1.0978 |  | 0.047931 |  | 0.086642 |  | 0.80503 |  | 0.33021 |  | 0.42643 |  | 0.014393 |  | 0.16143 |  | 0.20331 |  | 0.16001 |  | 0.047684
Within-run %CV |  |  | 11.79 |  | 0.51529 |  | 0.94029 |  | 8.4625 |  | 3.7753 |  | 5.5292 |  | 0.17937 |  | 1.8937 |  | 2.4036 |  | 1.8491 |  | 0.51294
QC 2 | Nominal Conc (µg/L) | Run 1 |  | Run 2 |  | Run 3 |  | Run 4 |  | Run 5 |  | Run 6 |  | Run 7 |  | Run 8 |  | Run 9 |  | Run 10 |  | Run 11
|  | 2019-07-26 |  | 2019-07-29 |  | 2019-07-30 |  | 2019-07-31 |  | 2019-08-01 |  | 2019-08-02 |  | 2019-08-05 |  | 2019-08-06 |  | 2019-08-07 |  | 2019-08-08 |  | 2019-08-09
Rep 1 | 50 | 40.9128 | 40.5391 | 45.7875 | 46.5783 | 43.0226 | 43.2172 | 48.765 | 48.5516 | 43.3907 | 43.6114 | 43.7987 | 43.4218 | 44.1299 | 43.4885 | 42.0768 | 41.9906 | 44.7747 | 45.1614 | 45.1547 | 45.6701 | 46.6639 | 46.762
|  | 40.1654 |  | 47.369 |  | 43.4118 |  | 48.3382 |  | 43.8321 |  | 43.045 |  | 42.8472 |  | 41.9043 |  | 45.5481 |  | 46.1855 |  | 46.8601
Rep 2 |  | 42.4194 | 42.344 | 47.3139 | 47.2519 | 44.2181 | 44.0019 | 46.321 | 45.7282 | 43.3789 | 43.4513 | 41.4232 | 41.4596 | 40.5611 | 40.7657 | 42.618 | 42.6355 | 46.8419 | 46.2855 | 45.5151 | 45.2043 | 46.3982 | 46.8186
|  | 42.2687 |  | 47.19 |  | 43.7857 |  | 45.1353 |  | 43.5238 |  | 41.4959 |  | 40.9703 |  | 42.653 |  | 45.729 |  | 44.8935 |  | 47.2389
Mean |  |  | 41.4416 |  | 46.9151 |  | 43.6095 |  | 47.1399 |  | 43.5314 |  | 42.4407 |  | 42.1271 |  | 42.3131 |  | 45.7234 |  | 45.4372 |  | 46.7903
Std Dev |  |  | 1.2763 |  | 0.47635 |  | 0.55488 |  | 1.9965 |  | 0.11317 |  | 1.3875 |  | 1.9253 |  | 0.45604 |  | 0.79484 |  | 0.32937 |  | 0.039997
Within-run %CV |  |  | 3.0797 |  | 1.0153 |  | 1.2724 |  | 4.2352 |  | 0.25997 |  | 3.2694 |  | 4.5702 |  | 1.0778 |  | 1.7384 |  | 0.72489 |  | 0.085482
QC 3 | Nominal Conc (µg/L) | Run 1 |  | Run 2 |  | Run 3 |  | Run 4 |  | Run 5 |  | Run 6 |  | Run 7 |  | Run 8 |  | Run 9 |  | Run 10 |  | Run 11
|  | 2019-07-26 |  | 2019-07-29 |  | 2019-07-30 |  | 2019-07-31 |  | 2019-08-01 |  | 2019-08-02 |  | 2019-08-05 |  | 2019-08-06 |  | 2019-08-07 |  | 2019-08-08 |  | 2019-08-09
Rep 1 | 140 | 127.4268 | 127.8553 | 139.946 | 138.1697 | 139.2225 | 140.3039 | 134.265 | 134.0726 | 133.9139 | 134.5746 | 116.9623 | 117.6422 | 133.8056 | 133.1481 | 136.8194 | 136.4824 | 127.2239 | 127.8136 | 127.7788 | 127.6614 | 133.4653 | 133.8374
|  | 128.2838 |  | 136.3934 |  | 141.3852 |  | 133.8803 |  | 135.2352 |  | 118.3222 |  | 132.4907 |  | 136.1454 |  | 128.4032 |  | 127.544 |  | 134.2094
Rep 2 |  | 128.1218 | 127.0217 | 140.9979 | 140.9863 | 142.534 | 143.8506 | 130.2199 | 131.0835 | 131.4737 | 131.558 | 119.241 | 118.938 | 124.0613 | 126.553 | 136.2646 | 134.4391 | 133.4806 | 133.6442 | 128.6532 | 127.2027 | 135.3156 | 134.8097
|  | 125.9217 |  | 140.9746 |  | 145.1672 |  | 131.947 |  | 131.6422 |  | 118.6351 |  | 129.0446 |  | 132.6136 |  | 133.8079 |  | 125.7521 |  | 134.3037
Mean |  |  | 127.4385 |  | 139.578 |  | 142.0772 |  | 132.5781 |  | 133.0663 |  | 118.2901 |  | 129.8505 |  | 135.4608 |  | 130.7289 |  | 127.432 |  | 134.3235
Std Dev |  |  | 0.58945 |  | 1.9916 |  | 2.5079 |  | 2.1137 |  | 2.1331 |  | 0.91628 |  | 4.6635 |  | 1.4449 |  | 4.1229 |  | 0.32439 |  | 0.68752
Within-run %CV |  |  | 0.46254 |  | 1.4269 |  | 1.7652 |  | 1.5943 |  | 1.603 |  | 0.77461 |  | 3.5914 |  | 1.0666 |  | 3.1538 |  | 0.25456 |  | 0.51184
*values outside acceptable criteria and removed from calculations
Nominal Conc (µg/L) | Grand Mean (ug/L) | Std Dev | Bias (%) | Accuracy (%) | Between -run %CV
10 | 8.796 | 0.65941 | -12.0397 | 87.9603 | 7.4966
50 | 44.3154 | 2.2288 | -11.3692 | 88.6308 | 5.0294
140 | 131.8931 | 6.5064 | -5.7907 | 94.2093 | 4.9331
Linearity Data | - Cocaine
Nominal Conc (µg/L) | Run 1 | 2019-07-26 |  | Run 2 | 2019-07-29 |  | Run 3 | 2019-07-30 |  | Run 4 | 2019-07-31 |  | Run 5 | 2019-08-01 |  | Run 6 | 2019-08-02
| Peak Area |  | Ratio Coc/IS | Peak Area |  | Ratio Coc/IS | Peak Area |  | Ratio Coc/IS | Peak Area |  | Ratio Coc/IS | Peak Area |  | Ratio Coc/IS | Peak Area |  | Ratio Coc/IS
| Coc | Coc-d8 |  | Coc | Coc-d8 |  | Coc | Coc-d8 |  | Coc | Coc-d8 |  | Coc | Coc-d8 |  | Coc | Coc-d8
2.5 | 10670.1741 | 66628.6236 | 0.16014 | 15720.1354 | 96276.9769 | 0.16328 | 17016.4709 | 107710.388 | 0.15798 | 14834.963 | 93548.7091 | 0.15858 | 15420.7069 | 101531.2076 | 0.15188 | 12037.3976 | 71323.7026 | 0.16877
5 | 26647.7997 | 83627.9844 | 0.31865 | 30050.5157 | 94933.0485 | 0.31654 | 31402.0322 | 100416.8826 | 0.31272 | 28261.0082 | 93486.1438 | 0.3023 | 29855.113 | 99270.7127 | 0.30074 | 21754.2742 | 70866.3815 | 0.30698
25 | 133897.7337 | 83848.082 | 1.5969 | 161878.8742 | 98989.7177 | 1.6353 | 157347.1748 | 102425.6728 | 1.5362 | 153614.9592 | 99217.4462 | 1.5483 | 151882.019 | 99797.0897 | 1.5219 | 115713.0518 | 72225.2025 | 1.6021
50 | 275745.5139 | 83735.1402 | 3.2931 | 340683.3899 | 103399.3544 | 3.2948 | 338050.2905 | 103873.538 | 3.2544 | 323493.7827 | 99644.1907 | 3.2465 | 297862.9706 | 97997.4156 | 3.0395 | 263157.0021 | 76571.3031 | 3.4368
75 | 443022.3416 | 87986.4215 | 5.0351 | 532697.9694 | 103237.8254 | 5.1599 | 547935.8641 | 108500.7518 | 5.0501 | 515617.561 | 101208.772 | 5.0946 | 490152.2378 | 96738.3656 | 5.0668 | 442281.6603 | 75396.65 | 5.8661
100 | 584956.9946 | 89117.9283 | 6.5639 | 747275.0036 | 109710.4292 | 6.8113 | 739350.4931 | 108748.6874 | 6.7987 | 621373.4212 | 90681.5986 | 6.8523 | 684413.4595 | 105868.119 | 6.4648 | 610560.6007 | 80063.5736 | 7.6259
150 | 908045.6835 | 89397.9907 | 10.1573 | 1051911.3096 | 102904.3392 | 10.2222 | 1073841.2493 | 107625.7659 | 9.9775 | 1027916.025 | 100098.0927 | 10.2691 | 993740.285 | 103902.0433 | 9.5642 | 932430.651 | 80914.9652 | 11.5236
200 | 1203145.9315 | 89824.2058 | 13.3945 | 1488589.9165 | 106564.0361 | 13.969 | 1452952.9894 | 106590.6215 | 13.6312 | 1422661.1722 | 104731.0358 | 13.584 | 1405993.2831 | 103373.3147 | 13.6011 | 1296211.4844 | 83028.4351 | 15.6117
Nominal Conc (µg/L) | Run 7 | 2019-08-05 |  | Run 8 | 2019-08-06 |  | Run 9 | 2019-08-07 |  | Run 10 | 2019-08-08 |  | Run 11 | 2019-08-09 |  | Average Peak Area Ratio
| Peak Area |  | Ratio Coc/IS | Peak Area |  | Ratio Coc/IS | Peak Area |  | Ratio Coc/IS | Peak Area |  | Ratio Coc/IS | Peak Area |  | Ratio Coc/IS
| Coc | Coc-d8 |  | Coc | Coc-d8 |  | Coc | Coc-d8 |  | Coc | Coc-d8 |  | Coc | Coc-d8
2.5 | 10651.2668 | 77203.1074 | 0.13796 | 14181.041 | 94122.1477 | 0.15067 | 14804.6964 | 92899.424 | 0.15936 | 21012.6637 | 111885.6354 | 0.1878 | 21701.4364 | 123272.8179 | 0.17604 | 0.16113
5 | 20511.8216 | 79874.7112 | 0.2568 | 26637.0496 | 88835.185 | 0.29985 | 25659.5327 | 93680.7007 | 0.2739 | 40088.4051 | 119341.5224 | 0.33591 | 40737.3783 | 123430.6294 | 0.33004 | 0.30495
25 | 115416.8117 | 77061.8585 | 1.4977 | 149233.893 | 95368.8692 | 1.5648 | 143991.8199 | 98429.3236 | 1.4629 | 208632.4646 | 122059.4939 | 1.7093 | 202614.7287 | 122409.0406 | 1.6552 | 1.5755
50 | 275269.6585 | 85021.165 | 3.2377 | 313937.2386 | 96115.0611 | 3.2663 | 299758.1996 | 97919.7069 | 3.0613 | 450162.57 | 127370.9616 | 3.5343 | 441023.9661 | 128727.1647 | 3.426 | 3.281
75 | 393469.7219 | 73724.42 | 5.337 | 488219.1153 | 90692.5655 | 5.3832 | 515140.9653 | 98668.8698 | 5.2209 | 692016.3103 | 121123.0645 | 5.7133 | 678318.0681 | 130339.1302 | 5.2043 | 5.2847
100 | 627649.1943 | 90745.2353 | 6.9166 | 720263.9631 | 100154.1824 | 7.1916 | 653616.2983 | 97222.1727 | 6.7229 | 932748.2559 | 123033.6111 | 7.5812 | 912977.723 | 131596.4296 | 6.9377 | 6.9515
150 | 845759.5588 | 81479.3936 | 10.38 | 1095974.6412 | 107602.5672 | 10.1854 | 981484.9111 | 100030.4035 | 9.8119 | 1274009.7181 | 122993.6864 | 10.3583 | 1372986.57 | 132717.1317 | 10.3452 | 10.2541
200 | 1134534.3841 | 79492.1624 | 14.2723 | 1486533.9891 | 104220.4159 | 14.2634 | 1436781.4276 | 99641.9044 | 14.4194 | 1821824.1653 | 123605.717 | 14.739 | 1902292.1347 | 134830.8967 | 14.1087 | 14.1449

## Precision Uncertainty
Precision Data |  | - Cocaine
Runs |  | Nominal Conc (NC) (µg/L) | Mean | Std Dev (SD) | Degrees of Freedom (DF) | Q = SD2 * DF | Pooled Variance Vp = (∑Q)/∑DF) | Pooled Std Dev (SDp) = √(Vp) | Case Sample reps (rcs) | Standard Uncertainty SU = SDp/√D | *Relative SU ur(Pre)=SU/NC
Run 1 | 2019-07-26 | 10 | 9.3114 | 1.0978 | 1 | 1.2052 | 0.20449 | 0.45221 | 2 | 0.31976 | 0.031976
Run 2 | 2019-07-29 |  | 9.3018 | 0.047931 | 1 | 0.0022974
Run 3 | 2019-07-30 |  | 9.2144 | 0.086642 | 1 | 0.0075069
Run 4 | 2019-07-31 |  | 9.5129 | 0.80503 | 1 | 0.64807 | [5 empty] | Precision DF PreDF = ∑DF=22
Run 5 | 2019-08-01 |  | 8.7466 | 0.33021 | 1 | 0.10904
Run 6 | 2019-08-02 |  | 7.7124 | 0.42643 | 1 | 0.18185
Run 7 | 2019-08-05 |  | 8.0242 | 0.014393 | 1 | 0.00020716
Run 8 | 2019-08-06 |  | 8.5247 | 0.16143 | 1 | 0.02606
Run 9 | 2019-08-07 |  | 8.4582 | 0.20331 | 1 | 0.041333
Run 10 | 2019-08-08 |  | 8.6535 | 0.16001 | 1 | 0.025603
Run 11 | 2019-08-09 |  | 9.2962 | 0.047684 | 1 | 0.0022737
|  |  |  | SUM | 11 | 2.2494
Runs |  | Nominal Conc (NC) (µg/L) | Mean | Std Dev (SD) | Degrees of Freedom (DF) | Q = SD2 * DF | Pooled Variance Vp = (∑Q)/∑DF) | Pooled Std Dev (SDp) = √(Vp) | Case Sample reps (rcs) | Standard Uncertainty SU = SDp/√D | *Relative SU ur(Pre)=SU/NC
Run 1 | 2019-07-26 | 50 | 41.4416 | 1.2763 | 1 | 1.6289 | 1.1586 | 1.0764 | 2 | 0.76111 | 0.015222
Run 2 | 2019-07-29 |  | 46.9151 | 0.47635 | 1 | 0.22691
Run 3 | 2019-07-30 |  | 43.6095 | 0.55488 | 1 | 0.30789
Run 4 | 2019-07-31 |  | 47.1399 | 1.9965 | 1 | 3.9859
Run 5 | 2019-08-01 |  | 43.5314 | 0.11317 | 1 | 0.012807
Run 6 | 2019-08-02 |  | 42.4407 | 1.3875 | 1 | 1.9253
Run 7 | 2019-08-05 |  | 42.1271 | 1.9253 | 1 | 3.7068
Run 8 | 2019-08-06 |  | 42.3131 | 0.45604 | 1 | 0.20797
Run 9 | 2019-08-07 |  | 45.7234 | 0.79484 | 1 | 0.63177
Run 10 | 2019-08-08 |  | 45.4372 | 0.32937 | 1 | 0.10848
Run 11 | 2019-08-09 |  | 46.7903 | 0.039997 | 1 | 0.0015998
|  |  |  | SUM | 11 | 12.7443
Runs |  | Nominal Conc (NC) (µg/L) | Mean | Std Dev (SD) | Degrees of Freedom (DF) | Q = SD2 * DF | Pooled Variance Vp = (∑Q)/∑DF) | Pooled Std Dev (SDp) = √(Vp) | Case Sample reps (rcs) | Standard Uncertainty SU = SDp/√D | *Relative SU ur(Pre)=SU/NC
Run 1 | 2019-07-26 | 140 | 127.4385 | 0.58945 | 1 | 0.34745 | 5.6248 | 2.3717 | 2 | 1.677 | 0.011979
Run 2 | 2019-07-29 |  | 139.578 | 1.9916 | 1 | 3.9666
Run 3 | 2019-07-30 |  | 142.0772 | 2.5079 | 1 | 6.2896
Run 4 | 2019-07-31 |  | 132.5781 | 2.1137 | 1 | 4.4676
Run 5 | 2019-08-01 |  | 133.0663 | 2.1331 | 1 | 4.55
Run 6 | 2019-08-02 |  | 118.2901 | 0.91628 | 1 | 0.83957
Run 7 | 2019-08-05 |  | 129.8505 | 4.6635 | 1 | 21.7481
Run 8 | 2019-08-06 |  | 135.4608 | 1.4449 | 1 | 2.0876
Run 9 | 2019-08-07 |  | 130.7289 | 4.1229 | 1 | 16.9982
Run 10 | 2019-08-08 |  | 127.432 | 0.32439 | 1 | 0.10523
Run 11 | 2019-08-09 |  | 134.3235 | 0.68752 | 1 | 0.47268
|  |  |  | SUM | 11 | 61.8726
*Use the relative standard uncertainity of precision (ur(Pre)) which has a nominal concentration (NC) closest to the mean of the analytical results obtained for the case sample.

## Std Solution & Spike Prep
Standard Solution Preparation and Curve Spiking |  |  |  |  | - Cocaine
Autopipettes (pip-10, 25, 50, 100, 250, 1000) | [8 empty] | Volumetric flask (Vol-10) (10 mL = 10000 µL) |  |  |  | Uncertainty of Matrix Volume
Parameters |  | Values | [6 empty] | Parameters |  | Values
Vol (µL) |  | 10 | 25 | 50 | 100 | 250 | 1000 |  | Vol (µL) |  | 10000 |  | Formula for the relative uncertainty of Matrix Volume = ur(Mtx) = |  |  |  | ur(pip-250(200))2
Accuracy (AC (± µL)) |  | 0.09 | 0.3 | 0.3 | 0.3 | 2.6 | 5 |  | Accuracy (AC (± µL)) |  | 25
Ref Cov factor (rCF) |  | 2 | 2 | 2 | 2 | 2 | 2 |  | Ref Cov factor (rCF) |  | 1.7321
Uncertatinty AcF = AC/rCF |  | 0.045 | 0.15 | 0.15 | 0.15 | 1.3 | 2.5 |  | Uncertatinty AcF = AC/rCF |  | 14.4338
ur(pip)2 |  | 2.025e-05 | 3.6e-05 | 9e-06 | 2.25e-06 | 2.704e-05 | 6.25e-06 |  | ur(Vol-10)2 |  | 2.0833e-06
RS Uncertainty of pipette |  | 0.0045 | 0.006 | 0.003 | 0.0015 | 0.0052 | 0.0025 |  | RS Uncertainty of Vol(10) |  | 0.0014434
Reference Std |  | Coc |  | Reference Int Std |  | Coc-d8 | [6 empty] | Result for Sample Vol (200 uL)
Parameters |  | Values |  | Parameters |  | Values | [6 empty] | Parameters |  | Values
CRM conc (mg/mL) |  | 1 |  | CRM conc (mg/mL) |  | 0.1004 | [6 empty] | ur(pip-250)2 |  | 2.704e-05
(µg/L) |  | 1000000 |  | (µg/L) |  | 100400 | [6 empty] | ur(Mtx)2 |  | 2.704e-05
Accuracy (AC (± µg/L)) |  | 5000 |  | Accuracy (AC (± µg/L)) |  | 1000 | [6 empty] | RS Uncertainty of Matrix Vol |  | 0.0052
Ref Cov factor (rCF) |  | 2 |  | Ref Cov factor (rCF) |  | 2
Uncertatinty AcF = AC/rCF |  | 2500 |  | Uncertatinty AcF = AC/rCF |  | 500
ur(Ref)2 |  | 6.25e-06 |  | ur(ISRef)2 |  | 2.4801e-05
RS Uncertainty of Ref |  | 0.0025 |  | RS Uncertainty of ISRef |  | 0.0049801
Coc-CAL Working Solution Prep | [10 empty] | Coc-IS Working Solution Prep
Formula for the relative uncertainty of CAL-WS1 Prep = ur(CAL-WS1) = | [5 empty] | ur(Ref)2 + Σ[ur(pip-X)2] |  |  |  |  | Formula for the relative uncertainty of IS-WS1 Prep = ur(IS-WS1) = |  |  |  |  | ur(ISRef)2 + Σ[ur(pip-X)2]
Formula for the relative uncertainty of CAL-WS2 Prep = ur(CAL-WS2) = | [5 empty] | ur(Ref)2+ Σ[ur(pip-X)2] |  |  |  |  | Formula for the relative uncertainty of ISSP Prep = ur(ISSP) = |  |  |  |  | ur(IS-WS1)2 + Σ[ur(pip-X)2] + ur(Vol-10)2
(from CRM)
Formula for the relative uncertainty of CAL-WS2 Prep = ur(CAL-WS2) = | [5 empty] | ur(CAL-WS1)2+ Σ[ur(pip-X)2]
(serial dilution of CAL-WS1)
Formula for the relative uncertainty of CAL-WS3 Prep = ur(CAL-WS3) = | [5 empty] | ur(Ref)2+ Σ[ur(pip-X)2]
Formula for the relative uncertainty of CAL-WS4 Prep = ur(CAL-WS4) = | [5 empty] | ur(Ref)2+ Σ[ur(pip-X)2]
Formula for the relative uncertainty of CALSP-X Prep = ur(CALSP-X) = | [5 empty] | Σ[ur(CAL-WSX)2] + Σ[ur(pip-X)2]
Aliquotting of drug and solvent
|  |  | (CRM) | (serial dilution)
Pipette / Volumetric Use |  | CAL-WS1 | CAL-WS2 | CAL-WS2 | CAL-WS3 | CAL-WS4 |  |  | CALSP-1 | CALSP-2 | CALSP-3 | CALSP-4 | CALSP-5 | CALSP-6 | CALSP-7 | CALSP-8 | IS-WS1 | ISSP
Pip-10
Pip-25 | [17 empty] | 1
Pip-50 | bar: CALSP-1
Pip-100 | bar: CALSP-2; CALSP-5; CALSP-8
Pip-250 |  | 1 |  | 1 | [6 empty] | 2 | 3 | 4 | 1 | 2 | 1
Pip-1000 |  | 2 |  | 3 |  |  |  |  | 3 | 3 | 3 | 2 | 2 | 3 | 2 | 2
Vol-10 | [17 empty] | 1
[8 empty] | CAL-WS1 | [8 empty] | IS-WS1 | bar: CALSP-6 to CALSP-8
[8 empty] | CAL-WS2 (CRM) | [8 empty] | CRM | 1
CAL-WS2 (sd) | bar: CALSP-1 to CALSP-5
CAL-WS3
CAL-WS4
CRM
Result for CAL-WS1 |  |  |  | Result for CAL-WS2 (CRM) |  |  |  | Result for CAL-WS2 (serial dilution) | CALSP-4: Result for CAL-WS3; CALSP-8: Result for CAL-WS4
Parameters |  | Values |  | Parameters |  | Values |  | Parameters | CALSP-2: Values; CALSP-4: Parameters; CALSP-6: Values; CALSP-8: Parameters | Values
ur(Ref)2 |  | 6.25e-06 |  | ur(Ref)2 |  | 6.25e-06 |  | ur(CAL-WS1)2 | CALSP-2: 4.579e-05; CALSP-4: ur(Ref)2; CALSP-6: 6.25e-06; CALSP-8: ur(Ref)2 | 6.25e-06
ur(pip-10)2 |  | 0 |  | ur(pip-10)2 |  | 0 |  | ur(pip-10)2 | CALSP-2: 0; CALSP-4: ur(pip-10)2; CALSP-6: 0; CALSP-8: ur(pip-10)2 | 0
ur(pip-25)2 |  | 0 |  | ur(pip-25)2 |  | 0 |  | ur(pip-25)2 | CALSP-2: 0; CALSP-4: ur(pip-25)2; CALSP-6: 0; CALSP-8: ur(pip-25)2 | 0
ur(pip-50)2 |  | 0 |  | ur(pip-50)2 |  | 0 |  | ur(pip-50)2 | CALSP-2: 0; CALSP-4: ur(pip-50)2; CALSP-6: 0; CALSP-8: ur(pip-50)2 | 0
ur(pip-100)2 |  | 0 |  | ur(pip-100)2 |  | 0 |  | ur(pip-100)2 | CALSP-2: 0; CALSP-4: ur(pip-100)2; CALSP-6: 0; CALSP-8: ur(pip-100)2 | 0
ur(pip-250)2 |  | 2.704e-05 |  | ur(pip-250)2 |  | 0 |  | ur(pip-250)2 | CALSP-2: 2.704e-05; CALSP-4: ur(pip-250)2; CALSP-6: 0; CALSP-8: ur(pip-250)2 | 0
ur(pip-1000)2 |  | 1.25e-05 |  | ur(pip-1000)2 |  | 0 |  | ur(pip-1000)2 | CALSP-2: 1.875e-05; CALSP-4: ur(pip-1000)2; CALSP-6: 0; CALSP-8: ur(pip-1000)2 | 0
ur(Vol-10)2 |  | 0 |  | ur(Vol-10)2 |  | 0 |  | ur(Vol-10)2 | CALSP-2: 0; CALSP-4: ur(Vol-10)2; CALSP-6: 0; CALSP-8: ur(Vol-10)2 | 0
ur(CAL-WS1)2 |  | 4.579e-05 |  | ur(CAL-WS2)2 |  | 6.25e-06 |  | ur(CAL-WS2)2 | CALSP-2: 9.158e-05; CALSP-4: ur(CAL-WS3)2; CALSP-6: 6.25e-06; CALSP-8: ur(CAL-WS4)2 | 6.25e-06
RS Uncertainty of CAL-WS1 |  | 0.0067668 |  | RS Uncertainty of CAL-WS2 |  | 0.0025 |  | RS Uncertainty of CAL-WS2 | CALSP-2: 0.0095697; CALSP-4: RS Uncertainty of CAL-WS3; CALSP-6: 0.0025; CALSP-8: RS Uncertainty of CAL-WS4 | 0.0025
Result for CALSP-1 |  |  |  | Result for CALSP-2 |  |  |  | Result for CALSP-3 | CALSP-4: Result for CALSP-4
Parameters |  | Values |  | Parameters |  | Values |  | Parameters | CALSP-2: Values; CALSP-4: Parameters; CALSP-6: Values
ur(CAL-WS1)2 |  | 0 |  | ur(CAL-WS1)2 |  | 0 |  | ur(CAL-WS1)2 | CALSP-2: 0; CALSP-4: ur(CAL-WS1)2; CALSP-6: 0
ur(CAL-WS2)2 (CRM) |  | 0 |  | ur(CAL-WS2)2 (CRM) |  | 0 |  | ur(CAL-WS2)2 (CRM) | CALSP-2: 0; CALSP-4: ur(CAL-WS2)2 (CRM); CALSP-6: 0
ur(CAL-WS2)2 (serial diln) |  | 9.158e-05 |  | ur(CAL-WS2)2 (serial diln) |  | 9.158e-05 |  | ur(CAL-WS2)2 (serial diln) | CALSP-2: 9.158e-05; CALSP-4: ur(CAL-WS2)2 (serial diln); CALSP-6: 9.158e-05
ur(CAL-WS3)2 |  | 0 |  | ur(CAL-WS3)2 |  | 0 |  | ur(CAL-WS3)2 | CALSP-2: 0; CALSP-4: ur(CAL-WS3)2; CALSP-6: 0
ur(CAL-WS4)2 |  | 0 |  | ur(CAL-WS4)2 |  | 0 |  | ur(CAL-WS4)2 | CALSP-2: 0; CALSP-4: ur(CAL-WS4)2; CALSP-6: 0
ur(Ref)2 |  | 0 |  | ur(Ref)2 |  | 0 |  | ur(Ref)2 | CALSP-2: 0; CALSP-4: ur(Ref)2; CALSP-6: 0
ur(pip-10)2 |  | 0 |  | ur(pip-10)2 |  | 0 |  | ur(pip-10)2 | CALSP-2: 0; CALSP-4: ur(pip-10)2; CALSP-6: 0
ur(pip-25)2 |  | 0 |  | ur(pip-25)2 |  | 0 |  | ur(pip-25)2 | CALSP-2: 0; CALSP-4: ur(pip-25)2; CALSP-6: 0
ur(pip-50)2 |  | 9e-06 |  | ur(pip-50)2 |  | 0 |  | ur(pip-50)2 | CALSP-2: 0; CALSP-4: ur(pip-50)2; CALSP-6: 0
ur(pip-100)2 |  | 0 |  | ur(pip-100)2 |  | 2.25e-06 |  | ur(pip-100)2 | CALSP-2: 0; CALSP-4: ur(pip-100)2; CALSP-6: 0
ur(pip-250)2 |  | 0 |  | ur(pip-250)2 |  | 0 |  | ur(pip-250)2 | CALSP-2: 5.408e-05; CALSP-4: ur(pip-250)2; CALSP-6: 8.112e-05
ur(pip-1000)2 |  | 1.875e-05 |  | ur(pip-1000)2 |  | 1.875e-05 |  | ur(pip-1000)2 | CALSP-2: 1.875e-05; CALSP-4: ur(pip-1000)2; CALSP-6: 1.25e-05
ur(Vol-10)2 |  | 0 |  | ur(Vol-10)2 |  | 0 |  | ur(Vol-10)2 | CALSP-2: 0; CALSP-4: ur(Vol-10)2; CALSP-6: 0
ur(CALSP-1)2 |  | 0.00011933 |  | ur(CALSP-2)2 |  | 0.00011258 |  | ur(CALSP-3)2 | CALSP-2: 0.00016441; CALSP-4: ur(CALSP-4)2; CALSP-6: 0.0001852
RS Uncertainty of CALSP-1 |  | 0.010924 |  | RS Uncertainty of CALSP-2 |  | 0.01061 |  | RS Uncertainty of CALSP-3 | CALSP-2: 0.012822; CALSP-4: RS Uncertainty of CALSP-4; CALSP-6: 0.013609
Result for CALSP-5 |  |  |  | Result for CALSP-6 |  |  |  | Result for CALSP-7 | CALSP-4: Result for CALSP-8
Parameters |  | Values |  | Parameters |  | Values |  | Parameters | CALSP-2: Values; CALSP-4: Parameters; CALSP-6: Values
ur(CAL-WS1)2 |  | 0 |  | ur(CAL-WS1)2 |  | 4.579e-05 |  | ur(CAL-WS1)2 | CALSP-2: 4.579e-05; CALSP-4: ur(CAL-WS1)2; CALSP-6: 4.579e-05
ur(CAL-WS2)2 (CRM) |  | 0 |  | ur(CAL-WS2)2 (CRM) |  | 0 |  | ur(CAL-WS2)2 (CRM) | CALSP-2: 0; CALSP-4: ur(CAL-WS2)2 (CRM); CALSP-6: 0
ur(CAL-WS2)2 (serial diln) |  | 9.158e-05 |  | ur(CAL-WS2)2 (serial diln) |  | 0 |  | ur(CAL-WS2)2 (serial diln) | CALSP-2: 0; CALSP-4: ur(CAL-WS2)2 (serial diln); CALSP-6: 0
ur(CAL-WS3)2 |  | 0 |  | ur(CAL-WS3)2 |  | 0 |  | ur(CAL-WS3)2 | CALSP-2: 0; CALSP-4: ur(CAL-WS3)2; CALSP-6: 0
ur(CAL-WS4)2 |  | 0 |  | ur(CAL-WS4)2 |  | 0 |  | ur(CAL-WS4)2 | CALSP-2: 0; CALSP-4: ur(CAL-WS4)2; CALSP-6: 0
ur(Ref)2 |  | 0 |  | ur(Ref)2 |  | 0 |  | ur(Ref)2 | CALSP-2: 0; CALSP-4: ur(Ref)2; CALSP-6: 0
ur(pip-10)2 |  | 0 |  | ur(pip-10)2 |  | 0 |  | ur(pip-10)2 | CALSP-2: 0; CALSP-4: ur(pip-10)2; CALSP-6: 0
ur(pip-25)2 |  | 0 |  | ur(pip-25)2 |  | 0 |  | ur(pip-25)2 | CALSP-2: 0; CALSP-4: ur(pip-25)2; CALSP-6: 0
ur(pip-50)2 |  | 0 |  | ur(pip-50)2 |  | 0 |  | ur(pip-50)2 | CALSP-2: 0; CALSP-4: ur(pip-50)2; CALSP-6: 0
ur(pip-100)2 |  | 2.25e-06 |  | ur(pip-100)2 |  | 0 |  | ur(pip-100)2 | CALSP-2: 0; CALSP-4: ur(pip-100)2; CALSP-6: 2.25e-06
ur(pip-250)2 |  | 0.00010816 |  | ur(pip-250)2 |  | 2.704e-05 |  | ur(pip-250)2 | CALSP-2: 5.408e-05; CALSP-4: ur(pip-250)2; CALSP-6: 2.704e-05
ur(pip-1000)2 |  | 1.25e-05 |  | ur(pip-1000)2 |  | 1.875e-05 |  | ur(pip-1000)2 | CALSP-2: 1.25e-05; CALSP-4: ur(pip-1000)2; CALSP-6: 1.25e-05
ur(Vol-10)2 |  | 0 |  | ur(Vol-10)2 |  | 0 |  | ur(Vol-10)2 | CALSP-2: 0; CALSP-4: ur(Vol-10)2; CALSP-6: 0
ur(CALSP-5)2 |  | 0.00021449 |  | ur(CALSP-6)2 |  | 9.158e-05 |  | ur(CALSP-7)2 | CALSP-2: 0.00011237; CALSP-4: ur(CALSP-8)2; CALSP-6: 8.758e-05
RS Uncertainty of CALSP-5 |  | 0.014645 |  | RS Uncertainty of CALSP-6 |  | 0.0095697 |  | RS Uncertainty of CALSP-7 | CALSP-2: 0.0106; CALSP-4: RS Uncertainty of CALSP-8; CALSP-6: 0.0093584
Result for IS-WS1 |  |  |  | Result for ISSP
Parameters |  | Values |  | Parameters |  | Values
ur(ISRef)2 |  | 2.4801e-05 |  | ur(IS-WS1)2 |  | 0
ur(pip-10)2 |  | 0 |  | ur(ISRef)2 |  | 2.4801e-05
ur(pip-25)2 |  | 0 |  | ur(pip-10)2 |  | 0
ur(pip-50)2 |  | 0 |  | ur(pip-25)2 |  | 3.6e-05
ur(pip-100)2 |  | 0 |  | ur(pip-50)2 |  | 0
ur(pip-250)2 |  | 0 |  | ur(pip-100)2 |  | 0
ur(pip-1000)2 |  | 0 |  | ur(pip-250)2 |  | 0
ur(Vol-10)2 |  | 0 |  | ur(pip-1000)2 |  | 0
ur(IS-WS1)2 |  | 2.4801e-05 |  | ur(Vol-10)2 |  | 2.0833e-06
RS Uncertainty of IS-WS1 |  | 0.0049801 |  | ur(ISSP)2 |  | 6.2885e-05
|  |  |  | RS Uncertainty of ISSP |  | 0.00793
Cocaine Calibration Curve Spiking
Formula for the relative uncertainty of calibration curve spiking = ur(CCSp) =
| ur(CCSp) = | ur(CALSP-1)2 + ur(CALSP-2)2 + ur(CALSP-3)2 + ur(CALSP-4)2 + ur(CALSP-5)2 + ur(CALSP-6)2 + ur(CALSP-7)2 + ur(CALSP-8)2 + ur(Cal spk vol)2 + ur(ISSP)2 + ur(IS spk vol)2 + ur(matrix vol)2
| ur(CCSp) = | ur(CALSP-1)2 + ur(CALSP-2)2 + ur(CALSP-3)2 + ur(CALSP-4)2 + ur(CALSP-5)2 + ur(CALSP-6)2 + ur(CALSP-7)2 + ur(CALSP-8)2 + 8*[ur(Pip-10)2] + ur(ISSP)2 + 8*[ur(Pip-10)2] + 8*[ur(Mtx)2]
Result for calibration curve spiking
Parameters |  |  | Values
ur(CALSP-1)2 |  |  | 0.00011933
ur(CALSP-2)2 |  |  | 0.00011258
ur(CALSP-3)2 |  |  | 0.00016441
ur(CALSP-4)2 |  |  | 0.0001852
ur(CALSP-5)2 |  |  | 0.00021449
ur(CALSP-6)2 |  |  | 9.158e-05
ur(CALSP-7)2 |  |  | 0.00011237
ur(CALSP-8)2 |  |  | 8.758e-05
ur(pip-10)2 x 8 |  |  | 0.000162
ur(ISSP)2 |  | excluded
ur(pip-10)2 x 8 |  |  | 0.000162
ur(Vol)2 x 8 |  |  | 0.00021632
ur(CCSp)2 |  |  | 0.0016279
RS Uncertainty of Cal Curve Spiking |  |  | 0.040347
UM for each calibrators 1-10 - first each cal working solution and subsequent spiking solution, then spiking of each cal.
Also includes IS prep/spiking and cal blood volume.
Based on diluting with pipettes only (not volumetrics)

## Calibration and Measurement
Calibration Uncertainty |  | - Cocaine
Nominal Conc (µg/L) | Ratio (Coc/Coc-d8)
x | y
| Run 1 | Run 2 | Run 3 | Run 4 | Run 5 | Run 6 | Run 7 | Run 8 | Run 9 | Run 10 | Run 11
| 2019-07-26 | 2019-07-29 | 2019-07-30 | 2019-07-31 | 2019-08-01 | 2019-08-02 | 2019-08-05 | 2019-08-06 | 2019-08-07 | 2019-08-08 | 2019-08-09
2.5 | 0.16014 | 0.16328 | 0.15798 | 0.15858 | 0.15188 | 0.16877 | 0.13796 | 0.15067 | 0.15936 | 0.1878 | 0.17604
5 | 0.31865 | 0.31654 | 0.31272 | 0.3023 | 0.30074 | 0.30698 | 0.2568 | 0.29985 | 0.2739 | 0.33591 | 0.33004
25 | 1.5969 | 1.6353 | 1.5362 | 1.5483 | 1.5219 | 1.6021 | 1.4977 | 1.5648 | 1.4629 | 1.7093 | 1.6552
50 | 3.2931 | 3.2948 | 3.2544 | 3.2465 | 3.0395 | 3.4368 | 3.2377 | 3.2663 | 3.0613 | 3.5343 | 3.426
75 | 5.0351 | 5.1599 | 5.0501 | 5.0946 | 5.0668 | 5.8661 | 5.337 | 5.3832 | 5.2209 | 5.7133 | 5.2043
100 | 6.5639 | 6.8113 | 6.7987 | 6.8523 | 6.4648 | 7.6259 | 6.9166 | 7.1916 | 6.7229 | 7.5812 | 6.9377
150 | 10.1573 | 10.2222 | 9.9775 | 10.2691 | 9.5642 | 11.5236 | 10.38 | 10.1854 | 9.8119 | 10.3583 | 10.3452
200 | 13.3945 | 13.969 | 13.6312 | 13.584 | 13.6011 | 15.6117 | 14.2723 | 14.2634 | 14.4194 | 14.739 | 14.1087
SEw | 0.0013019 | 0.0018827 | 0.0023382 | 0.0028295 | 0.002894 | 0.0061108 | 0.005181 | 0.0036935 | 0.0057835 | 0.0038973 | 0.001861 | Total
(n-1) | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 7 | 77
(n-1)SEw2 | 1.1864e-05 | 2.4811e-05 | 3.8272e-05 | 5.6041e-05 | 5.8625e-05 | 0.0002614 | 0.0001879 | 9.5491e-05 | 0.00023414 | 0.00010632 | 2.4243e-05 | 0.0010991
Largest SE for | 2019-07-26: Run 6; 2019-07-29: 2019-08-02; 2019-08-07: Regression Parameters
Run 6 | Nominal Conc (µg/L) | Ratio (Coc/Coc-d8) | Weights | xw | yw | wx2 |  | Total Readings (n) | Calibration Df (CurDF) | Intercept (b0) | Slope (b1) | Std Error (SEw) | Weighted X Mean (x̅w)
| x | y | w = 1/x^2
| 2.5 | 0.16877 | 0.16 | 0.4 | 0.027003 | 1 |  | 8 | 6 | -0.02595 | 0.073639 | 0.0061108 | 0.086875
| 5 | 0.30698 | 0.04 | 0.2 | 0.012279 | 1
| 25 | 1.6021 | 0.0016 | 0.04 | 0.0025634 | 1 |  | Case Sample
2019-08-02 | 50 | 3.4368 | 0.0004 | 0.02 | 0.0013747 | 1 |  | Mean Concentration (µg/L) (x̅cs) | Peak Area Rotio (ycs) | No. Case Sample Replicates (rcs) |  | Mean Concentration (µg/L) (x̅cs) | wcs = 1/xcs2
| 75 | 5.8661 | 0.00017778 | 0.013333 | 0.0010429 | 1
| 100 | 7.6259 | 0.0001 | 0.01 | 0.00076259 | 1
| 150 | 11.5236 | 4.4444e-05 | 0.0066667 | 0.00051216 | 1 |  | 10 | 0.71044 | 2 |  | 10 | 0.01
| 200 | 15.6117 | 2.5e-05 | 0.005 | 0.00039029 | 1 |  | 50 | 3.656 |  |  | 50 | 0.0004
| 2019-07-31: x̅w = (∑xw)/n; 2019-08-01: y̅w = (∑xw)/n; 2019-08-02: ∑wx2; 2019-08-06: 140; 2019-08-07: 10.2836 | 140 | 5.102e-05
|  |  |  | 0.086875 | 0.0057411 | 8
Weighted Regression Analysis
OVERALL FIT | 2019-08-08: Pooled Standard Error (SEp)
Multiple R | 2019-07-26: 0.99634; 2019-07-30: AIC; 2019-07-31: -81.8646
R Square | 2019-07-26: 0.99269; 2019-07-30: AICc; 2019-07-31: -79.4646; 2019-08-08: SEw(p) =; 2019-08-09: (n1-1)SE12 + …. + (nj-1)SEj2
Adjusted R Square | 2019-07-26: 0.99165; 2019-07-30: BSC; 2019-07-31: -81.7851; 2019-08-09: (n1-1) + …. + (nj-1)
Standard Error | 2019-07-26: 0.0061108
Observations | 2019-07-26: 8
[10 empty] | = | 0.0010991
ANOVA | 2019-07-31: Alpha; 2019-08-01: 0.05; 2019-08-09: 77
| df | SS | MS | F | p-value | sig
Regression | 1 | 0.030437 | 0.030437 | 815.0921 | 1.2227e-07 | yes |  |  |  | = | 0.0037781
Residual | 2019-07-26: 6; 2019-07-29: 0.00022405; 2019-07-30: 3.7342e-05
Total | 2019-07-26: 7; 2019-07-29: 0.030661; 2019-08-08: SEw(p) is used as a reliable estimate for SEw
| coeff | std err | t stat | p-value | lower | upper | vif
Intercept | 2019-07-26: -0.02595; 2019-07-29: 0.016218; 2019-07-30: -1.6; 2019-07-31: 0.16071; 2019-08-01: -0.065634; 2019-08-02: 0.013735
x | 0.073639 | 0.0025793 | 28.5498 | 1.2227e-07 | 0.067328 | 0.079951 | 1
Calibration Curve
Formula for the uncertainty of the calibration curve = u(Cur) = | 2019-07-31: SEw(p); 2019-08-01: 1; 2019-08-02: +; 2019-08-05: 1; 2019-08-06: +; 2019-08-07: (ycs - y̅w)2
|  |  |  | b1 | wcs |  | n |  | b12[∑wx2-n(x̅w)2]
u(Cur) = | 2019-07-31: A; 2019-08-01: B; 2019-08-02: +; 2019-08-05: C; 2019-08-06: +; 2019-08-07: D
E
Relative Standard Uncertainty = ur(Cur) = | 2019-08-01: u(Cur)/x̅cs
Result for calibration curve at Conc = 10 µg/L | 2019-07-31: Result for calibration curve at Conc = 50 µg/L; 2019-08-06: Result for calibration curve at Conc = 140 µg/L
Parameters | Formula | Values |  | Parameters | Formula | Values |  | Parameters | Formula | Values
A | SEw(p)/b1 | 0.051306 |  | A | SEw(p)/b1 | 0.051306 |  | A | SEw(p)/b1 | 0.051306
B | 1/wcs | 100 |  | B | 1/wcs | 2500 |  | B | 1/wcs | 19600
C | 1/n | 0.125 |  | C | 1/n | 0.125 |  | C | 1/n | 0.125
D (x̅cs = 10) | (ycs - y̅w)2 | 0.49661 |  | D (x̅cs = 50) | (ycs - y̅w)2 | 13.3245 |  | D (x̅cs = 140) | (ycs - y̅w)2 | 105.6334
E | b12[∑wx2-n(x̅w)2] | 0.043055 |  | E | b12[∑wx2-n(x̅w)2] | 0.043055 |  | E | b12[∑wx2-n(x̅w)2] | 0.043055
u(Cur10) | 2019-07-29: 0.54214; 2019-07-31: u(Cur50); 2019-08-02: 2.7195; 2019-08-06: u(Cur140); 2019-08-08: 7.6191
ur(Cur10)^2 | 2019-07-29: 0.0029392; 2019-07-31: ur(Cur50)^2; 2019-08-02: 0.0029582; 2019-08-06: ur(Cur140)^2; 2019-08-08: 0.0029618
ur(Cur10) | u(Cur10)/Con10 | 0.054214 |  | ur(Cur50) | u(Cur50)/Con50 | 0.05439 |  | ur(Cur140) | u(Cur140)/Con140 | 0.054422
Analytical Measurement
Formula for the relative uncertainty of analytical measurement = ur(x̅cs) = | 2019-08-01: ur(CCSp)2 + ur(Cur)2
Result for analytical measurement; Conc = 10 µg/L | 2019-07-31: Result for analytical measurement; Conc = 50 µg/L; 2019-08-06: Result for analytical measurement; Conc = 140 µg/L
Parameters | 2019-07-29: Values; 2019-07-31: Parameters; 2019-08-02: Values; 2019-08-06: Parameters; 2019-08-08: Values
ur(CCSp)2 | 2019-07-29: 0.0016279; 2019-07-31: ur(CCSp)2; 2019-08-02: 0.0016279; 2019-08-06: ur(CCSp)2; 2019-08-08: 0.0016279
ur(Cur10)^2 | 2019-07-29: 0.0029392; 2019-07-31: ur(Cur50)^2; 2019-08-02: 0.0029582; 2019-08-06: ur(Cur140)^2; 2019-08-08: 0.0029618
ur(x̅cs)2 | 2019-07-29: 0.004567; 2019-07-31: ur(x̅cs)2; 2019-08-02: 0.0045861; 2019-08-06: ur(x̅cs)2; 2019-08-08: 0.0045896
RS Uncertainty of analytical measurement | 2019-07-29: 0.06758; 2019-07-31: RS Uncertainty of analytical measurement; 2019-08-02: 0.067721; 2019-08-06: RS Uncertainty of analytical measurement; 2019-08-08: 0.067747
run1 | 2019-08-07: run2
Weighted Regression Analysis | 2019-08-08: Weighted Regression Analysis
OVERALL FIT | 2019-08-07: OVERALL FIT
Multiple R | 2019-07-26: 0.99979; 2019-07-30: AIC; 2019-07-31: -106.6047; 2019-08-07: Multiple R; 2019-08-08: 0.99959 | AIC | -100.7026
R Square | 2019-07-26: 0.99959; 2019-07-30: AICc; 2019-07-31: -104.2047; 2019-08-07: R Square; 2019-08-08: 0.99917 | AICc | -98.3026
Adjusted R Square | 2019-07-26: 0.99953; 2019-07-30: BSC; 2019-07-31: -106.5253; 2019-08-07: Adjusted R Square; 2019-08-08: 0.99906 | BSC | -100.6232
Standard Error | 2019-07-26: 0.0013019; 2019-08-07: Standard Error; 2019-08-08: 0.0018827
Observations | 2019-07-26: 8; 2019-08-07: Observations; 2019-08-08: 8
ANOVA | 2019-07-31: Alpha; 2019-08-01: 0.05; 2019-08-07: ANOVA |  | Alpha | 0.05
| df | SS | MS | F | p-value | sig |  |  |  | df | SS | MS | F | p-value | sig
Regression | 1 | 0.024638 | 0.024638 | 14536.6253 | 2.195e-11 | yes |  |  | Regression | 1 | 0.025706 | 0.025706 | 7252.6819 | 1.7655e-10 | yes
Residual | 2019-07-26: 6; 2019-07-29: 1.0169e-05; 2019-07-30: 1.6949e-06; 2019-08-07: Residual; 2019-08-08: 6; 2019-08-09: 2.1266e-05 | 3.5444e-06
Total | 2019-07-26: 7; 2019-07-29: 0.024648; 2019-08-07: Total; 2019-08-08: 7; 2019-08-09: 0.025728
| coeff | std err | t stat | p-value | lower | upper | vif |  |  | coeff | std err | t stat | p-value | lower | upper | vif
Intercept | -0.0072491 | 0.0034552 | -2.098 | 0.080696 | -0.015704 | 0.0012054 |  |  | Intercept | -0.009444 | 0.0049966 | -1.8901 | 0.10764 | -0.02167 | 0.0027822
x | 0.066253 | 0.00054951 | 120.5679 | 2.195e-11 | 0.064908 | 0.067598 | 1 |  | x | 0.067675 | 0.00079465 | 85.1627 | 1.7655e-10 | 0.06573 | 0.069619 | 1
run3 | 2019-08-07: run4
Weighted Regression Analysis | 2019-08-07: Weighted Regression Analysis
OVERALL FIT | 2019-08-07: OVERALL FIT
Multiple R | 2019-07-26: 0.99933; 2019-07-30: AIC; 2019-07-31: -97.2351; 2019-08-07: Multiple R; 2019-08-08: 0.99904 | AIC | -94.1841
R Square | 2019-07-26: 0.99867; 2019-07-30: AICc; 2019-07-31: -94.8351; 2019-08-07: R Square; 2019-08-08: 0.99807 | AICc | -91.7841
Adjusted R Square | 2019-07-26: 0.99848; 2019-07-30: BSC; 2019-07-31: -97.1557; 2019-08-07: Adjusted R Square; 2019-08-08: 0.9978 | BSC | -94.1047
Standard Error | 2019-07-26: 0.0023382; 2019-08-07: Standard Error; 2019-08-08: 0.0028295
Observations | 2019-07-26: 8; 2019-08-07: Observations; 2019-08-08: 8
ANOVA | 2019-07-31: Alpha; 2019-08-01: 0.05; 2019-08-07: ANOVA |  | Alpha | 0.05
| df | SS | MS | F | p-value | sig |  |  |  | df | SS | MS | F | p-value | sig
Regression | 1 | 0.024583 | 0.024583 | 4496.2969 | 7.3998e-10 | yes |  |  | Regression | 1 | 0.024885 | 0.024885 | 3108.3844 | 2.2362e-09 | yes
Residual | 2019-07-26: 6; 2019-07-29: 3.2804e-05; 2019-07-30: 5.4674e-06; 2019-08-07: Residual; 2019-08-08: 6; 2019-08-09: 4.8035e-05 | 8.0059e-06
Total | 2019-07-26: 7; 2019-07-29: 0.024616; 2019-08-07: Total; 2019-08-08: 7; 2019-08-09: 0.024933
| coeff | std err | t stat | p-value | lower | upper | vif |  |  | coeff | std err | t stat | p-value | lower | upper | vif
Intercept | -0.010315 | 0.0062057 | -1.6622 | 0.14754 | -0.0255 | 0.00487 |  |  | Intercept | -0.013092 | 0.0075094 | -1.7434 | 0.13188 | -0.031467 | 0.0052827
x | 0.06618 | 0.00098695 | 67.0544 | 7.3998e-10 | 0.063765 | 0.068595 | 1 |  | x | 0.066585 | 0.0011943 | 55.7529 | 2.2362e-09 | 0.063663 | 0.069508 | 1
run5 | 2019-08-07: run6
Weighted Regression Analysis | 2019-08-07: Weighted Regression Analysis
OVERALL FIT | 2019-08-07: OVERALL FIT
Multiple R | 2019-07-26: 0.99892; 2019-07-30: AIC; 2019-07-31: -93.8235; 2019-08-07: Multiple R; 2019-08-08: 0.99634 | AIC | -81.8646
R Square | 2019-07-26: 0.99784; 2019-07-30: AICc; 2019-07-31: -91.4235; 2019-08-07: R Square; 2019-08-08: 0.99269 | AICc | -79.4646
Adjusted R Square | 2019-07-26: 0.99754; 2019-07-30: BSC; 2019-07-31: -93.7441; 2019-08-07: Adjusted R Square; 2019-08-08: 0.99165 | BSC | -81.7851
Standard Error | 2019-07-26: 0.002894; 2019-08-07: Standard Error; 2019-08-08: 0.0061108
Observations | 2019-07-26: 8; 2019-08-07: Observations; 2019-08-08: 8
ANOVA | 2019-07-31: Alpha; 2019-08-01: 0.05; 2019-08-07: ANOVA |  | Alpha | 0.05
| df | SS | MS | F | p-value | sig |  |  |  | df | SS | MS | F | p-value | sig
Regression | 1 | 0.023253 | 0.023253 | 2776.5057 | 3.1358e-09 | yes |  |  | Regression | 1 | 0.030437 | 0.030437 | 815.0921 | 1.2227e-07 | yes
Residual | 2019-07-26: 6; 2019-07-29: 5.025e-05; 2019-07-30: 8.375e-06; 2019-08-07: Residual; 2019-08-08: 6; 2019-08-09: 0.00022405 | 3.7342e-05
Total | 2019-07-26: 7; 2019-07-29: 0.023304; 2019-08-07: Total; 2019-08-08: 7; 2019-08-09: 0.030661
| coeff | std err | t stat | p-value | lower | upper | vif |  |  | coeff | std err | t stat | p-value | lower | upper | vif
Intercept | -0.012056 | 0.0076806 | -1.5697 | 0.16754 | -0.03085 | 0.0067375 |  |  | Intercept | -0.02595 | 0.016218 | -1.6 | 0.16071 | -0.065634 | 0.013735
x | 0.064365 | 0.0012215 | 52.6926 | 3.1358e-09 | 0.061376 | 0.067354 | 1 |  | x | 0.073639 | 0.0025793 | 28.5498 | 1.2227e-07 | 0.067328 | 0.079951 | 1
run7 | 2019-08-07: run8
Weighted Regression Analysis | 2019-08-07: Weighted Regression Analysis
OVERALL FIT | 2019-08-07: OVERALL FIT
Multiple R | 2019-07-26: 0.99688; 2019-07-30: AIC; 2019-07-31: -84.5055; 2019-08-07: Multiple R; 2019-08-08: 0.99846 | AIC | -89.9206
R Square | 2019-07-26: 0.99377; 2019-07-30: AICc; 2019-07-31: -82.1055; 2019-08-07: R Square; 2019-08-08: 0.99692 | AICc | -87.5206
Adjusted R Square | 2019-07-26: 0.99289; 2019-07-30: BSC; 2019-07-31: -84.4261; 2019-08-07: Adjusted R Square; 2019-08-08: 0.99648 | BSC | -89.8411
Standard Error | 2019-07-26: 0.005181; 2019-08-07: Standard Error; 2019-08-08: 0.0036935
Observations | 2019-07-26: 8; 2019-08-07: Observations; 2019-08-08: 8
ANOVA | 2019-07-31: Alpha; 2019-08-01: 0.05; 2019-08-07: ANOVA |  | Alpha | 0.05
| df | SS | MS | F | p-value | sig |  |  |  | df | SS | MS | F | p-value | sig
Regression | 1 | 0.025709 | 0.025709 | 957.765 | 7.558e-08 | yes |  |  | Regression | 1 | 0.026469 | 0.026469 | 1940.3444 | 9.1653e-09 | yes
Residual | 2019-07-26: 6; 2019-07-29: 0.00016106; 2019-07-30: 2.6843e-05; 2019-08-07: Residual; 2019-08-08: 6; 2019-08-09: 8.185e-05 | 1.3642e-05
Total | 2019-07-26: 7; 2019-07-29: 0.02587; 2019-08-07: Total; 2019-08-08: 7; 2019-08-09: 0.026551
| coeff | std err | t stat | p-value | lower | upper | vif |  |  | coeff | std err | t stat | p-value | lower | upper | vif
Intercept | -0.042207 | 0.013751 | -3.0695 | 0.021957 | -0.075853 | -0.0085603 |  |  | Intercept | -0.026345 | 0.0098025 | -2.6875 | 0.036171 | -0.05033 | -0.0023587
x | 0.067679 | 0.0021869 | 30.9478 | 7.558e-08 | 0.062328 | 0.07303 | 1 |  | x | 0.068672 | 0.001559 | 44.0493 | 9.1653e-09 | 0.064857 | 0.072486 | 1
run9 | 2019-08-07: run10
Weighted Regression Analysis | 2019-08-07: Weighted Regression Analysis
OVERALL FIT | 2019-08-07: OVERALL FIT
Multiple R | 2019-07-26: 0.99584; 2019-07-30: AIC; 2019-07-31: -82.7455; 2019-08-07: Multiple R; 2019-08-08: 0.99843 | AIC | -89.0609
R Square | 2019-07-26: 0.9917; 2019-07-30: AICc; 2019-07-31: -80.3455; 2019-08-07: R Square; 2019-08-08: 0.99687 | AICc | -86.6609
Adjusted R Square | 2019-07-26: 0.99052; 2019-07-30: BSC; 2019-07-31: -82.6661; 2019-08-07: Adjusted R Square; 2019-08-08: 0.99642 | BSC | -88.9814
Standard Error | 2019-07-26: 0.0057835; 2019-08-07: Standard Error; 2019-08-08: 0.0038973
Observations | 2019-07-26: 8; 2019-08-07: Observations; 2019-08-08: 8
ANOVA | 2019-07-31: Alpha; 2019-08-01: 0.05; 2019-08-07: ANOVA |  | Alpha | 0.05
| df | SS | MS | F | p-value | sig |  |  |  | df | SS | MS | F | p-value | sig
Regression | 1 | 0.023987 | 0.023987 | 717.1433 | 1.7906e-07 | yes |  |  | Regression | 1 | 0.029015 | 0.029015 | 1910.2202 | 9.6046e-09 | yes
Residual | 2019-07-26: 6; 2019-07-29: 0.00020069; 2019-07-30: 3.3448e-05; 2019-08-07: Residual; 2019-08-08: 6; 2019-08-09: 9.1135e-05 | 1.5189e-05
Total | 2019-07-26: 7; 2019-07-29: 0.024188; 2019-08-07: Total; 2019-08-08: 7; 2019-08-09: 0.029106
| coeff | std err | t stat | p-value | lower | upper | vif |  |  | coeff | std err | t stat | p-value | lower | upper | vif
Intercept | -0.014914 | 0.015349 | -0.97164 | 0.36876 | -0.052473 | 0.022645 |  |  | Intercept | 0.0013222 | 0.010344 | 0.12782 | 0.90246 | -0.023988 | 0.026632
x | 0.065373 | 0.0024412 | 26.7795 | 1.7906e-07 | 0.0594 | 0.071346 | 1 |  | x | 0.071898 | 0.001645 | 43.7061 | 9.6046e-09 | 0.067873 | 0.075923 | 1
run11
Weighted Regression Analysis
OVERALL FIT
Multiple R | 2019-07-26: 0.99961; 2019-07-30: AIC; 2019-07-31: -100.8878
R Square | 2019-07-26: 0.99921; 2019-07-30: AICc; 2019-07-31: -98.4878
Adjusted R Square | 2019-07-26: 0.9991; 2019-07-30: BSC; 2019-07-31: -100.8083
Standard Error | 2019-07-26: 0.001861
Observations | 2019-07-26: 8
ANOVA | 2019-07-31: Alpha; 2019-08-01: 0.05
| df | SS | MS | F | p-value | sig
Regression | 2019-07-26: 1; 2019-07-29: 0.026434; 2019-07-30: 0.026434; 2019-07-31: 7632.6527; 2019-08-01: 1.5149e-10; 2019-08-02: yes
Residual | 2019-07-26: 6; 2019-07-29: 2.078e-05; 2019-07-30: 3.4633e-06
Total | 2019-07-26: 7; 2019-07-29: 0.026455
| coeff | std err | t stat | p-value | lower | upper | vif
Intercept | 2019-07-26: 0.00061176; 2019-07-29: 0.0049391; 2019-07-30: 0.12386; 2019-07-31: 0.90547; 2019-08-01: -0.011474; 2019-08-02: 0.012697
x | 0.068626 | 0.00078551 | 87.3651 | 1.5149e-10 | 0.066704 | 0.070548 | 1

## Combined UM
Combined Uncertatinty |  |  | - Cocaine
Combined Uncertainty
|  |  | Formula for the combined uncertainty = comU = | Conc | ∑(Individual Uncertainty)2
|  |  | comU = | Conc | *ur(Pre)2 + ur(CCSp)2 + ur(Cur)2 + ur(Mtx)2
|  |  | comU = | Conc | *ur(Pre)2 + ur(x̅cs)2 + ur(Mtx)2
*Use the relative standard uncertainity of precision (ur(Pre)) which has a nominal concentration (NC) closest to the mean of the analytical results obtained for the case sample.
Eg. A sample concentration close to a NC of 10 will use ur(Pre) = 0.0319760976701291
Concentration = | 10 µg/L
| comU = | 10 | 0.0010225 | + | 0.004567 | + | 2.704e-05
| comU = | 10 | 0.0056165
| comU = | 0.74944
Concentration = | 50 µg/L
| comU = | 50 | 0.00023171 | + | 0.0045861 | + | 2.704e-05
| comU = | 50 | 0.0048449
| comU = | 3.4803
Concentration = | 140 µg/L
| comU = | 140 | 0.00014349 | + | 0.0045896 | + | 2.704e-05
| comU = | 140 | 0.0047602
| comU = | 9.6592

## Effective DF
Effetive Degrees of Freedom |  | - Cocaine
Effective degree of freedom (DFeff)
Formula for effective degrees of freedom = DFeff = |  |  | (Combined Uncertainty/Concentration)4
∑((Individual Uncertainty)4/(Individual DF))
|  | DFeff = | (comU/Conc)4
|  |  | ur(Pre)4 | + | ur(CCSp)4 | + | ur(Cur)4 | + | ur(Mtx)4
|  |  | PreDF |  | CCSpDF* |  | CurDF |  | MtxDF*
*The degrees of freedom of the standard solution and matrix are unknown and therefore CCSpDF and MtxDF = ∞ (infinity) which gives 0 when divided.
| Concentration = | 10 ug/L
|  | DFeff = | 3.1546e-05
|  |  | 1.0454e-06 | + | 0 | + | 8.6387e-06 | + | 0
|  |  | 11 |  |  |  | 6
|  | DFeff = | 3.1546e-05
1.5348e-06
|  | DFeff = | 20.5531
| Concentration = | 50 ug/L
|  | DFeff = | 2.3473e-05
|  |  | 5.3692e-08 | + | 0 | + | 8.7512e-06 | + | 0
|  |  | 11 |  |  |  | 6
|  | DFeff = | 2.3473e-05
1.4634e-06
|  | DFeff = | 16.0396
| Concentration = | 140 ug/L
|  | DFeff = | 2.2659e-05
|  |  | 2.0589e-08 | + | 0 | + | 8.7722e-06 | + | 0
|  |  | 11 |  |  |  | 6
|  | DFeff = | 2.2659e-05
1.4639e-06
|  | DFeff = | 15.4787

## Expanded UM
Expanded Uncertainty |  | - Cocaine
Expanded Uncertainty
|  | Formula for the expanded uncertainty = ExpU = | k x Combined Uncertainty
|  | ExpU = | 3*x comU
where k is the coverage factor for approximately 99.7% probability.
*Given that the effective degress of freedom DFeff (computations given in sheet='Effective DF') is large enough to assume a normal distribution, a coverage factor of k = 3 is used for a coverage probabilty of 99.7%.
Use k = 2 for 95%.
Concentration = 10 µg/L |  | Concentration = 50 µg/L |  |  | Concentration = 140 µg/L
ExpU = | 2.2483 | ExpU = | 10.4408 |  | ExpU = | 28.9775
ExpU% = | 22.4831 | ExpU% = | 20.8815 |  | ExpU% = | 20.6982
The reported value is Conc ± ExpU or Conc ± ExpU%, where the percentage of uncertainity (ExpU%) is given by
Percentage Expanded Uncertainty
|  |  | Formula for the percentage expanded uncertainty = ExpU% = | (ExpU/Conc) x 100